## PATCH_SETTINGS
offset | size | type | State | Field Name in Code | Value | description | Field Name in Requirements
0 | 1 | uint8 | N/A | CapsuleStateHFRFlag | 4 | Bit #n represents whether capsule state #n is HFR (1) or not (0).
1 | 1 | uint8 | N/A | FrameEncryptionEnable | 1 | Enable ('1') encryption of frames before eMMC storage
2 | 1 | uint8 | N/A | BLEEncryptionEnable | 1 | Enable ('1') encryption of BLE communication outside Production state (always disabled in production).
3 | 1 | uint8 | N/A | EngOpt | 0 | Engineering Optional Bit
4 | 4 | uint32 | N/A | MaxRTVSessionTime | 30000 | [msec] Maximum RTV session duration | MaxRTVSessionTime
8 | 4 | uint32 | N/A | MAXRTVTotalTime | 1800000 | [msec] Maximum accumulated duration of all RTV sessions | MAXRTVTotalTime
12 | 4 | uint32 | N/A | RTVConnectionTimeout | 45000 | [msec] Maximum time patch will wait for TCP Socket successful connection until reset of wireless module into sleep mode
16 | 4 | uint32 | N/A | FuelGaugeQueryTime | 55000 | [msec] Periodic cycle of power and thermal management
20 | 4 | uint32 | N/A | NominalBatteryCapacity | 2200 | [mA*h] Capacity of patch battery
24 | 4 | uint32 | N/A | BatteryThreshold | 90 | [%] Minimum remaining patch battery level to allow a procedure to start | BatteryThreshold
28 | 4 | uint32 | N/A | CapsuleWTCThreshold | 5000 | Maximum PillCam WTC value to allow a procedure to start (Default 5000) | CapsuleWTCThreshold
32 | 4 | uint32 | N/A | CapsuleSearchTimeout | 1800000 | [msec] Timeout for receiving the first PillCam frame in a procedure after the MODEM Rx is turned ON, before shutting down the patch. | CapsuleSearchTimeout
36 | 4 | uint32 | N/A | CapsuleSearchDelay | 6000 | [msec] Delay in Capsule Search state between patch turn on and MODEM RX turn on. | CapsuleSearchDelay
40 | 4 | uint32 | N/A | PSD_override_mode_psd_reg | 0 | For R&D debug usage only
44 | 4 | uint32 | N/A | DummyFrameTimeout | 5000 | [msec] Time from the last processed frame to generate a Dummy frame | t_dummy
48 | 4 | uint32 | N/A | ShutDownTemperatureThreshold | 480 | [0.1*oC] Patch shutdown thermal threshold | shdn_temp_th
52 | 4 | uint32 | N/A | ShutDownBatteryThreshold | 0 | [%] Patch shutdown battery threshold | shdn_bat_th
56 | 4 | uint32 | N/A | ShutDownVoltageThreshold | 2050 | [mV] Patch shutdown battery voltage threshold (Default 2050) | shdn_volt_th
60 | 4 | uint32 | N/A | LowPowerTemperatureThreshold | 470 | [0.1*oC] Patch low power mode thermal threshold(520) | lp_temp_th
64 | 4 | uint32 | N/A | LowPowerBatteryThreshold | 3 | [%] Patch low power mode battery threshold | lp_bat_th
68 | 4 | uint32 | N/A | ReducedPowerTemperatureThreshold | 460 | [0.1*oC] Patch reduced power mode thermal threshold(470) | rp_temp_th
72 | 4 | uint32 | N/A | ReducedPowerBatteryThreshold | 8 | [%] Patch reduced power mode battery threshold | rp_bat_th
76 | 4 | uint32 | N/A | TemperatureHysteresis | 9 | [0.1*oC] Patch temperature hysteresis | temp_hysteresis
80 | 2 | uint16 | N/A | RecordingIndicationType | 0 | 0: Timer based Green blink every 2 seconds. 1: Frame reception based pulse. | RecordingIndicationType
82 | 1 | uint8 | N/A | CMA_enable | 1 | 0: Disable CMA. 1: Enable CMA (Communication Management Algorithm)
83 | 1 | uint8 | N/A | PSD_override_mode_flag | 0 | For R&D debug usage only
84 | 4 | uint32 | N/A | wifi_retries | 3 | Number of WiFi transmissions retries
88 | 4 | uint32 | N/A | WiFi_timeout | 30 | [msec] WiFi respond timeout
92 | 4 | uint32 | N/A | ble_retries | 3 | Number of BLE transmissions retries
96 | 4 | uint32 | N/A | ble_timeout | 30000 | [msec] BLE respond timeout
100 | 4 | uint32 | N/A | BLEIdleTimeout | 300000 | [msec] BLE IDLE Timeout (to close IDLE BLE connections) | BLEIdleTimeout
104 | 4 | uint32 | N/A | t_enable_shutdown | 900000 | [msec] Time from beginning of Recording state when patch can still be turned off with a magnet. | t_enable_shutdown
108 | 4 | uint32 | N/A | ADCCalibrationInterval | 60000 | [msec] Time between consecutive calls to ADC re-callibration (per ADC)
112 | 4 | uint32 | N/A | FullProcedureTime | 43200000 | Intire procedure duration (12 Hours)
116 | 4 | uint32 | N/A | UploadTimout | 0 | 117 minutes +-3 upload timout from last success upload of file (burst of N frames) to S3
120 | 4 | uint32 | N/A | RESERVED3 | 0 | Reserved for future use
124 | 4 | uint32 | N/A | ProcedureStatesCount | 4 | number of procedure states: 1 to 10 | N
128 | 4 | uint32 | 0 | TriggerMaxDurationInState | 180000 | [msec] Maximum time in current procedure state before moving on to the next state.
132 | 4 | uint32 | 0 | TriggerMaxDurationGlobal | 0 | [msec] Maximum time in Recording state before moving on to the next procedure state.
136 | 4 | uint32 | 0 | NoFrameTimeout | 1500000 | [msec] Maximum time since last PillCam frame received before issuing EOP.
140 | 4 | uint32 | 0 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
144 | 1 | uint8 | 0 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
145 | 1 | uint8 | 0 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
146 | 1 | uint8 | 0 | AFRCapsuleStateLFR | 0 | Capsule state number that serves as LFR state for current procedure state.
147 | 1 | uint8 | 0 | AFRCapsuleStateHFR | 0 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
148 | 2 | uint16 | 0 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
150 | 2 | uint16 | 0 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
152 | 4 | uint32 | 0 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
156 | 4 | uint32 | 1 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
160 | 4 | uint32 | 1 | TriggerMaxDurationGlobal | 7200000 | [msec] Maximum time in Recording state before moving on to the next procedure state.
164 | 4 | uint32 | 1 | NoFrameTimeout | 1500000 | [msec] Maximum time since last PillCam frame received before issuing EOP.
168 | 4 | uint32 | 1 | TriggerCapsuleHFRCounter | 120000 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
172 | 1 | uint8 | 1 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
173 | 1 | uint8 | 1 | AFREnable | 1 | 1' = Enable AFR in current procedure state. '0' = disable
174 | 1 | uint8 | 1 | AFRCapsuleStateLFR | 1 | Capsule state number that serves as LFR state for current procedure state.
175 | 1 | uint8 | 1 | AFRCapsuleStateHFR | 2 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
176 | 2 | uint16 | 1 | AfrHIGH2LOWTH | 500 | AFR High to Low threshold. AFR score below this threshold sets LFR.
178 | 2 | uint16 | 1 | AfrLOW2HIGHTH | 500 | AFR Low to High threshold. AFR score above this threshold sets HFR.
180 | 4 | uint32 | 1 | AFRHoldTime | 1000 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
184 | 4 | uint32 | 2 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
188 | 4 | uint32 | 2 | TriggerMaxDurationGlobal | 43200000 | [msec] Maximum time in Recording state before moving on to the next procedure state.
192 | 4 | uint32 | 2 | NoFrameTimeout | 1500000 | [msec] Maximum time since last PillCam frame received before issuing EOP.
196 | 4 | uint32 | 2 | TriggerCapsuleHFRCounter | 120000 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
200 | 1 | uint8 | 2 | TriggerPatchBattery | 5 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
201 | 1 | uint8 | 2 | AFREnable | 1 | 1' = Enable AFR in current procedure state. '0' = disable
202 | 1 | uint8 | 2 | AFRCapsuleStateLFR | 1 | Capsule state number that serves as LFR state for current procedure state.
203 | 1 | uint8 | 2 | AFRCapsuleStateHFR | 2 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
204 | 2 | uint16 | 2 | AfrHIGH2LOWTH | 1600 | AFR High to Low threshold. AFR score below this threshold sets LFR.
206 | 2 | uint16 | 2 | AfrLOW2HIGHTH | 1600 | AFR Low to High threshold. AFR score above this threshold sets HFR.
208 | 4 | uint32 | 2 | AFRHoldTime | 1000 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
212 | 4 | uint32 | 3 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
216 | 4 | uint32 | 3 | TriggerMaxDurationGlobal | 43200001 | [msec] Maximum time in Recording state before moving on to the next procedure state.
220 | 4 | uint32 | 3 | NoFrameTimeout | 1500000 | [msec] Maximum time since last PillCam frame received before issuing EOP.
224 | 4 | uint32 | 3 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
228 | 1 | uint8 | 3 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
229 | 1 | uint8 | 3 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
230 | 1 | uint8 | 3 | AFRCapsuleStateLFR | 1 | Capsule state number that serves as LFR state for current procedure state.
231 | 1 | uint8 | 3 | AFRCapsuleStateHFR | 1 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
232 | 2 | uint16 | 3 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
234 | 2 | uint16 | 3 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
236 | 4 | uint32 | 3 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
240 | 4 | uint32 | 4 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
244 | 4 | uint32 | 4 | TriggerMaxDurationGlobal | 0 | [msec] Maximum time in Recording state before moving on to the next procedure state.
248 | 4 | uint32 | 4 | NoFrameTimeout | 0 | [msec] Maximum time since last PillCam frame received before issuing EOP.
252 | 4 | uint32 | 4 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
256 | 1 | uint8 | 4 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
257 | 1 | uint8 | 4 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
258 | 1 | uint8 | 4 | AFRCapsuleStateLFR | 0 | Capsule state number that serves as LFR state for current procedure state.
259 | 1 | uint8 | 4 | AFRCapsuleStateHFR | 0 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
260 | 2 | uint16 | 4 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
262 | 2 | uint16 | 4 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
264 | 4 | uint32 | 4 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
268 | 4 | uint32 | 5 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
272 | 4 | uint32 | 5 | TriggerMaxDurationGlobal | 0 | [msec] Maximum time in Recording state before moving on to the next procedure state.
276 | 4 | uint32 | 5 | NoFrameTimeout | 0 | [msec] Maximum time since last PillCam frame received before issuing EOP.
280 | 4 | uint32 | 5 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
284 | 1 | uint8 | 5 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
285 | 1 | uint8 | 5 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
286 | 1 | uint8 | 5 | AFRCapsuleStateLFR | 0 | Capsule state number that serves as LFR state for current procedure state.
287 | 1 | uint8 | 5 | AFRCapsuleStateHFR | 0 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
288 | 2 | uint16 | 5 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
290 | 2 | uint16 | 5 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
292 | 4 | uint32 | 5 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
296 | 4 | uint32 | 6 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
300 | 4 | uint32 | 6 | TriggerMaxDurationGlobal | 0 | [msec] Maximum time in Recording state before moving on to the next procedure state.
304 | 4 | uint32 | 6 | NoFrameTimeout | 0 | [msec] Maximum time since last PillCam frame received before issuing EOP.
308 | 4 | uint32 | 6 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
312 | 1 | uint8 | 6 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
313 | 1 | uint8 | 6 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
314 | 1 | uint8 | 6 | AFRCapsuleStateLFR | 0 | Capsule state number that serves as LFR state for current procedure state.
315 | 1 | uint8 | 6 | AFRCapsuleStateHFR | 0 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
316 | 2 | uint16 | 6 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
318 | 2 | uint16 | 6 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
320 | 4 | uint32 | 6 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
324 | 4 | uint32 | 7 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
328 | 4 | uint32 | 7 | TriggerMaxDurationGlobal | 0 | [msec] Maximum time in Recording state before moving on to the next procedure state.
332 | 4 | uint32 | 7 | NoFrameTimeout | 0 | [msec] Maximum time since last PillCam frame received before issuing EOP.
336 | 4 | uint32 | 7 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
340 | 1 | uint8 | 7 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
341 | 1 | uint8 | 7 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
342 | 1 | uint8 | 7 | AFRCapsuleStateLFR | 0 | Capsule state number that serves as LFR state for current procedure state.
343 | 1 | uint8 | 7 | AFRCapsuleStateHFR | 0 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
344 | 2 | uint16 | 7 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
346 | 2 | uint16 | 7 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
348 | 4 | uint32 | 7 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
352 | 4 | uint32 | 8 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
356 | 4 | uint32 | 8 | TriggerMaxDurationGlobal | 0 | [msec] Maximum time in Recording state before moving on to the next procedure state.
360 | 4 | uint32 | 8 | NoFrameTimeout | 0 | [msec] Maximum time since last PillCam frame received before issuing EOP.
364 | 4 | uint32 | 8 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
368 | 1 | uint8 | 8 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
369 | 1 | uint8 | 8 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
370 | 1 | uint8 | 8 | AFRCapsuleStateLFR | 0 | Capsule state number that serves as LFR state for current procedure state.
371 | 1 | uint8 | 8 | AFRCapsuleStateHFR | 0 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
372 | 2 | uint16 | 8 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
374 | 2 | uint16 | 8 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
376 | 4 | uint32 | 8 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.
380 | 4 | uint32 | 9 | TriggerMaxDurationInState | 0 | [msec] Maximum time in current procedure state before moving on to the next state.
384 | 4 | uint32 | 9 | TriggerMaxDurationGlobal | 0 | [msec] Maximum time in Recording state before moving on to the next procedure state.
388 | 4 | uint32 | 9 | NoFrameTimeout | 0 | [msec] Maximum time since last PillCam frame received before issuing EOP.
392 | 4 | uint32 | 9 | TriggerCapsuleHFRCounter | 0 | Number of HFR frames transmitted by PillCam threshold above which the patch transitions to the next procedure state.
396 | 1 | uint8 | 9 | TriggerPatchBattery | 0 | [%] Patch remaining battery level threshold below which the patch transitions to the next procedure state. 0 = disable
397 | 1 | uint8 | 9 | AFREnable | 0 | 1' = Enable AFR in current procedure state. '0' = disable
398 | 1 | uint8 | 9 | AFRCapsuleStateLFR | 0 | Capsule state number that serves as LFR state for current procedure state.
399 | 1 | uint8 | 9 | AFRCapsuleStateHFR | 0 | Capsule state number that serves as HFR state for current procedure state. Should be equals to "AFRCapsuleStateLFR" when AFR is disabled.
400 | 2 | uint16 | 9 | AfrHIGH2LOWTH | 0 | AFR High to Low threshold. AFR score below this threshold sets LFR.
402 | 2 | uint16 | 9 | AfrLOW2HIGHTH | 0 | AFR Low to High threshold. AFR score above this threshold sets HFR.
404 | 4 | uint32 | 9 | AFRHoldTime | 0 | [msec] Time to remain in HFR before going back to LFR, regardless of the AFR score.